GDP Receiver Quality Control checklist.

# Date:  _______________          GDP S/N : _____________________

## Front panel type & S/N: ____________________________________

Crew Chief:______________________   Intended Job #: _________

Note: Complete XMT checklist first, before starting receiver checklist.

- Open up battery compartment and check:

- Batteries in good condition &  Date on batteries: _____________________

- Measure each battery with multi-meter. Are they all similar voltages? If a battery has a lower voltage that others, it may require replacing. A matched set of batteries with the same date is preferred.

Battery voltages  _________,    _________,    _________,    _________.

- Battery terminals all in good condition and connected.

- Insulation tape covering terminals or areas that may short.

- Wiring harness in good condition with no damaged insulation

- External check of receiver:

- Are all screws installed? Check bottom panel of receiver, and then close battery compartment taking care not to pinch wiring.

- Receiver body and keyboard is clean?

- Binding posts on right hand side in reasonable condition?

- Grounding bar / jumper cable between case and analog ground in place?

- Ports on left-hand side are in reasonable condition and have dust covers?

- Binding posts in lid panel in reasonable condition?

- Analog meter ribbon cable installed and in reasonable condition?

- Black input lid cable in reasonable condition?

- Open lid and turn receiver on:

- Listen to beeps and hard drive turning / stopping. Does it sound normal?

- Note above EPROM or software version numbers

- Push continue button to initialize analog cards. Do they all pass?

- Check how many analog cards are installed. Is there enough for the job required? Record channel serial numbers below.

## Serial #’s 1:_____ 2:_____ 3:_____ 4:_____ 5:_____ 6:_____ 7:_____  8:_____

## 9:_____ 10:_____ 11:_____ 12:_____ 13:_____ 14:_____15:_____ 16:_____

- Check if NanoTEM cards are installed. Have them removed if not required.

- Choose RP/IP program and turn on all channels and continue to data acquisition mode. Remove ground bar and using a multi-meter measure resistance between COM and Case Gnd on I/O side panel. This should be infinite. A finite resistance (e.g. 10 M ohms) means there is a bad analog card.

- Measure resistance between each channel input (red & black) and Case Gnd. This should also be infinite.

- Measure resistance between COM and each channel input. The resistance should be approx. 4.9 to 5 M ohms. Any less or more means that channel has a problem and requires attention.

- Calibrates and systems checks

- Re-connect ground bar between COMM and case Gnd.

- Run internal calibrate in RP/IP mode with notch filter in. Run internal and external auto system tests and check data cache for any problems. Phase values should be 0.0. May be a little out on higher frequencies.

- Repeat above procedures with notch filter out.

- Repeat above 2 tests using CR program to test harmonic analysis mode.

- Go to main menu and select Sync option. Run sync check with XMT.

- Check coil calibrate cache for appropriate AMT / CSAMT coil calibrate serial numbers for the intended survey. Load correct file if values aren’t coils to be used.

- Sync Check:

- Return to main menu and run sync check in utility option

- Data transfer

- Connect serial port and dump data and checks to laptop.

- For GDP-32II use network cable to access data between receiver and laptop.

QC checks conducted by: _________________________________